Introduction

Welcome to the Imaging Basic Troubleshooting Guide. This guide is designed to assist users in diagnosing and resolving common issues encountered during image capturing procedures. Whether you are a seasoned professional or new to imaging technology, this guide provides practical solutions to ensure smooth and efficient operation.

Purpose of This Guide

The primary purpose of this guide is to empower users with the knowledge and tools necessary to troubleshoot and resolve imaging-related problems. By following the steps outlined here, users can minimize downtime and maintain the quality of their imaging processes.

Who Should Use This Guide

This guide is intended for:

- Dental professionals: Including office managers, dental assistants, and clinicians who will be using XVWeb for imaging purposes.

- IT support teams: Responsible for the technical setup and troubleshooting of XVWeb and associated hardware.

- Installation contacts: Individuals designated to oversee the installation process and coordinate with technicians.

| Issue | Possible causes | Workable Solution |
| --- | --- | --- |
| Sensor not connecting /error/ or nothing happening | - Old imaging software open in background - Capture window was minimized from previous user and “Done” was never selected - Wrong sensor chosen - Loose wire or extension used - Not installed or installed incorrectly - Tubehead too far from sensor | - Close all OTHER imaging software except Dentiray capture - Ensure there are no other capture windows open and have staff always hit “done” to save - Ensure correct sensor is chosen from device list - Remove extension and test sensor directly plugged into PC if possible - Ensure latest device drivers are installed - Ensure tubehead is as close to the sensor as possible - Submit support ticket |
| Missing Images-Newly acquired Xrays are missing/cannot find them | - Capture window was minimized from previous user and “Done” was never selected - Images are ghosted or labeled as retaken - Capture software was closed prior to images being forwarded to Web - Images had failed to forward to the web (offline, other network related problems) - Images were taken in wrong patient - Images were taken under a duplicate patient | - From workstation where images were taken ensure there are no other capture windows open and have staff always hit “Done” to save - Ensure images are NOT ghosted or Retake’s images and change labels. * If wrong layout was used Reorganize into new layout under Image Operations in Capture - From workstation where images were taken check the main screen of capture for forwarding errors (bottom box) see instructions on how to forward (Resubmitting Failed Forwards - XVCapture/DCV – Apteryx (zendesk.com)) - Utilizing the MRU button on the workstation where images were taken check other patients file OR using advanced search under view images you can search by date and even tooth number - Search the workstation where missing images were captured by ID only, then first name only, then last name only to locate any potiental duplicate patient records - Submit support ticket if images are still not showing (forwarding path may need to be adjusted) |
| Xrays are not showing in layout correctly (off center or at the bottom/wrong position) | - Layout was created in capture (progression editor) but NOT in Web - There is no tooth number associated to x ray - Tooth number association exists, but is incorrect - Web >Layout> View is set to incorrect layout - XVCapture View settings are not set correctly | - All layouts that are created in the progression editor of Capture ALSO need to be created in the Web for proper viewing (link) - If layout was used to take images ensure CORRECT layout was chosen, if so, then check the layout within progression editor and ensure teeth are associated with each tile/shot in the layout (this is sometimes overlooked) - If incorrect layout was used, use Layouts > Place to reorganize them into the correct layout - In Web choose Layouts> View and choose FMX - In XVCapture choose Image Operations > View Images> grouped anatomically Progressions Displayed as Grid Instead of Anatomical - XVCapture/DCV – Apteryx (zendesk.com) |
| Xrays are returning flipped or rotated (layout) | - Layout is missing sensor orientation - Driver settings need adjustment - Capture settings need adjustment | - Within Progression Editor find layout used> Choose edit layout> Edit each shot with correct sensor orientation (the software considers top of sensor the flat side opposing the wire) - Have office communicate with their IT to adjust driver settings (you can also submit support ticket for our team to assist them if absolutely needed) - Post capture filter needs to be created (Submit support ticket to ensure filter will not apply to multiple office locations) |
| Pan/Ceph is returning flipped or rotated | - Driver settings need adjustment - Capture settings need adjustment | - Have office communicate with their IT to adjust driver settings (you can also submit support ticket for our team to assist them if absolutely needed) - Post capture filter needs to be created , or Device Settings adjusted |
| Newly created layout is ONLY showing on one computer (XVCapture) | - Software is installed locally on each workstation | - Option A – Copy already created progression file from one workstation to other workstations how to copy layout to other workstations - Option B- recreate layout on each workstation (not recommended as this leaves room for error) |
| Image is too dark | - Overexposure or improper positioning, distance, or technique - Device driver settings may need adjustment - Software settings may need adjustment | - Decrease exposure time on generator, ensure proper placement of x ray and distance of generator cone. - Have office communicate with their IT to adjust driver settings (you can also submit support ticket for our team to assist them if needed) - Submit support ticket for image quality appointment |
| Image is too light | - Underexposure, improper positioning, distance, or technique - Device driver settings may need adjustment - Capture Software settings may need adjustment | - Increase exposure time on generator, ensure proper placement of x ray and distance of generator cone. - Have office communicate with their IT to adjust driver settings (you can also submit support ticket for our team to assist them if needed) - Submit support ticket for image quality appointment |
| Image is Blurry or Grainy | - Improper positioning, distance, or technique - Device driver settings may need adjustment - Software settings may need adjustment | - Ensure proper placement of device. Ensure proper distance and angle of generator cone. Hold exposure button until full cycle completes - Have office communicate with their IT to adjust driver settings (you can also submit support ticket for our team to assist them if needed) - Submit support ticket for image quality appointment |
| Brightness /Contrast will not save | - Office using enhancement features in wrong screen and NOT enhancement screen - User did not click Save after modifying B/C/G in Web - Office not saving enhancements when prompted - Modified image save settings need correction | - Brightness /contrast and gamma adjustments within Web can be changed and saved on a per-image basis. These can be modified within the XVCapture enhancement screen if preferred instead. - Ensure office is opening enhancement screen within XVCapture (using full screen view or capture screen prior to hitting done will not save brightness contrast adjustments) - Once enhancements are made and office chooses Save & Go Back ensure they are choosing yes to save option - Adjust saving of images in settings of capture Derived/Duplicate Images - XVCapture/DCV – Apteryx (zendesk.com) |
| Flip and rotate will not save or is saving as ghosted image | - User did not click “Save” if flip/rotate was applied through Web - XVCapture preferences set to save modified images as derivative images | - Best practice for flip/rotate to is do this in capture prior to hitting “Done” or select image from XVCapture collection screen and choose image operations> edit image information flip/rotate then okay - If images have already been saved, flips and rotates can be saved if done through Web via Transform tool - If office prefers enhancement screen for flip/rotate they will need to save & go back > click on stacked image and choose derived image as primary OR adjust modified image saving in preference settings |
| Measuring tool is not showing correct measurement | - Measuring tool was not calibrated or someone adjusted calibration - Office attempting to measure on Pan/Ceph or Plates | - Ensure measuring tool was calibrated correctly (see instructions) - Measuring tools in any imaging software are never recommended for pan/ceph or plates as there is distortion in the raw image *if plates are being used it is recommended that they calibrate EVERY time they use measuring tool (have them measure distance of plate itself for reference instead of coin) |